Calibrating the Prime Arm (FARO) – Measurement Session Notes

Presenter: Jeff Moss, Diverse Dimensions
Location: Lab setup before client visit

🛠️ Purpose of the Session

- Preparing for a measurement session at a tool builder’s site

- Using the FARO Prime Arm (6-axis arm, unlike the Quantum which has 7)

- Calibrating the arm before transport, due to limited setup space at the client site

🧰 Tools & Equipment

- FARO Prime Arm (6-axis, ideal for tooling inspection)

- PolyWorks 2020 software

- Calibration cone (setup ~1 joint-length away from the base of the arm)

- Various probes from MetrologyWorks:

  - 3mm ball probe with a long shank (great for reaching behind clamps)

  - 6mm ball probe (also used)

🧪 Calibration Setup Steps

- Set up arm and cone in lab at appropriate distance

- Launch PolyWorks and connect to the arm:

  - Ensure communication with device is established

  - Identify the correct device by serial number

- Select probe in PolyWorks:

  - Use probe manager to activate desired probe (e.g., 3mm)

- Initiate Compensation Procedure:

  - Navigate to Compensation tab and begin step-by-step instructions

  - System recognizes it’s a 6-axis arm and adjusts the calibration steps accordingly

🔧 Performing the Calibration

- Let the arm dangle naturally (due to counterbalance)

- Insert the probe into the calibration cone divots

- Start from the bottom and rotate upwards:

  - Insert probe gently into divot

  - Hold green button and rotate past 90°

  - Release button

  - Repeat this for 3 locations around the cone

- After 3 rotations, press the red button to complete calibration

✅ Calibration Outcome

- Successful calibration dialog appears

- Reported error: ~25 microns (acceptable by FARO specs)

- Probe is now ready for use

📌 Key Takeaways

- Calibrating in-lab saves time and ensures better conditions than onsite

- MetrologyWorks probes are reliable and versatile

- Proper calibration is crucial before any measurement session

- More advanced training available upon request

Capturing and Measuring Slot Features in PolyWorks (V14)

Location: Lab at Diverse Dimensions
Presented by: Jeff Moss, featuring work by Mike

🎯 Purpose of Video

- Demonstrate how to:

  - Capture slot components on sheet metal parts using hard probing

  - Handle a special dimensional requirement from a drawing

- Mike previously worked on two parts with slot features and an unusual dimension spec

🖥️ Software & Equipment

- PolyWorks V14

- Hard probe on FARO arm

- Prepped CAD models of parts

- Target: Slot component "dim 7-1" and special dimension "dim 8"

📏 Capturing a Slot in PolyWorks

- Open the model tree and right-click → Probe Measured

- Ensure constraining plane is already measured (important for slot capture)

- Probing technique:

  - Use finger to align equator of probe with trim edge of the slot

  - Take 3 points on first arc → press back button to compensate

  - Move to second arc → take 3 more points → press back button again

  - Ensure probe remains above the tangent point

- PolyWorks solves for the perfect center of the slot (visualized in 3D scene)

📐 Reviewing the Slot Data

- Zoom in on the slot to examine the feature

- Positioning shows variation in X-axis

- Close measurement view and open report tab in PolyWorks

📄 Drawing Integration

- Mike has drawing ballooned directly inside the PolyWorks report

- Avoids separate hand-marked printouts

- Dimension 8 (976 mm) is called out on drawing:

  - Points to top arc of slot, not the centerline

  - Requires capturing the arc feature itself for dimensional check

🔄 Creating an Arc Feature from the Slot

- Highlight slot feature "dim 7-1"

- Go to Circle tool in PolyWorks

- Method: From Objects

- With slot highlighted, choose Submethod: Slot Arcs

- Hit Create → software generates:

  - Circular arc from top portion of slot

  - Optionally another from bottom

- Use this arc to solve for dim 8 from the drawing

🧠 Tips & Insights

- PolyWorks includes many hidden or advanced tools not used every day

- If unsure how to approach a measurement:

  - Explore submethods and object types

  - Odds are the feature is supported

- If you find an unusual use case or want a video on it:

  - Contact Jeff — others likely want to see it too!

📬 Follow-Up

- For unique measurement scenarios, reach out with suggestions

- Future videos may cover rare or advanced techniques in PolyWorks

Measuring Cylindrical Features Accurately in PolyWorks

Location: Late night in the Diverse Dimensions lab
Presenter: Jeff Moss

🛠️ Context

- Working late to finish Tilt housings

- Demonstrating an important measurement issue with small cylinders

- Purpose: Show how to correctly form cylindrical features when limited data causes incorrect vector orientation

📌 Problem Overview

- When measuring small or narrow cylinders:

  - Not enough probe points = incorrect axis solution

  - Result: Cylinder may be formed with the wrong orientation

- Example: Intended cylinder is narrow along Y, but PolyWorks forms it incorrectly along X

🧪 Demonstration in PolyWorks

- Open pre-measured model in PolyWorks

- Feature in question: Datum D

  - Represents a full cylinder from two pivot points in the housing

- Initial Attempt (Standard Probe Method):

  - Took 4 points on one pivot, 4 on the other

  - PolyWorks formed a cylinder in the wrong direction

  - Result = unusable for accurate datum or axis measurement

✅ Correct Approach: Use Initial Circles

- PolyWorks Submethod: "Use Initial Circles"

- This mimics earlier methods from CAM2 (Faro software)

  - In that system: 2 circles define the base orientation, then additional points refine the cylinder

🔄 Steps:

- Select “Use Initial Circles” in probe menu

- Measure Circle 1 of 2:

  - Place 3 points on the first pivot ring

- Measure Circle 2 of 2:

  - Place 3 points on the second pivot ring

- Add additional probe points anywhere else as needed

- Hit compensate to finish measurement

🎯 Result:

- Cylinder is now correctly formed and oriented

- Appears accurately in model view and follows correct axis

💡 Takeaways & Tips

- If your cylinder appears in the wrong direction:

  - Try using the Initial Circles method

- Best for small, internal, or hard-to-access cylinders

- Only need:

  - 3 points per initial circle

  - Then as many other points as you want

- Helps ensure the axis vector is accurate

🧭 Applications

- Works for:

  - Internal cylindrical features (like pivot tubes)

  - External cylinders as well

- Useful when default probing doesn't interpret geometry correctly

Capturing a Plane Feature in PolyWorks (FARO Arm Tutorial)

Presenter: Diverse Dimensions team
Use Case: Reverse engineering die blocks

🧰 Project Context

- Working on a reverse engineering project

- Client brought in a number of die blocks

- Most features are prismatic in shape

  - Easily measured using planes, circles, cylinders, etc.

📐 Objective

- Demonstrate how to probe and capture a plane

- Though a basic feature, it's foundational to many models

🖥️ Procedure in PolyWorks

- Open PolyWorks model

- Navigate to the Feature toolbar at the top

- Click on "Create a Plane Feature"

  - Automatically enters probe mode

- Use the FARO arm to collect data points

✍️ Probing the Plane

- Minimum: 3 points required

- Can collect more points for higher accuracy

- In this demo: 4 points captured

- After capturing:

  - Press red button to lift arm

  - Hit "Compensate" to complete the feature

✅ Result

- Plane feature is now successfully captured in PolyWorks

- Ready to be used for alignment, feature referencing, or further modeling

🎯 Key Takeaway

- Plane probing is simple but essential for accurate geometry modeling in reverse engineering

- Knowing how to capture foundational shapes like planes ensures smoother downstream processes

Using the Tape Measure Macro in PolyWorks

Presenter: Derek from Diverse Dimensions
Location: Eastside Campus, Brighton, MI
Audience: Users wanting basic measurement tools in PolyWorks

🧰 Overview

- Demo of a useful macro for basic measurements in PolyWorks

- Ideal for users who:

  - Don’t need the full Inspector interface

  - Want a simplified measurement experience using FARO gear

🧭 Where to Find It

- Go to: PolyWorks Tech Support Zone

  - Navigate to: Macros → Inspector → Tape Measure

- Macro: “Tape Measure”

  - Enables basic measurements like:

    - Distances

    - Angles

  - Simplified UI for quicker use

🖼️ Visual Layout Tips

- This macro uses a custom “Tape Measure Visual Layout”

- ⚠️ Important:

  - Save your current layout before activating the macro

  - Look in the video description for a link on how to restore your layout

- Useful for:

  - Power users helping co-workers perform simpler tasks

  - Keeping workflows streamlined for less experienced users

🧪 Example Measurement Workflow

- Load the Tape Measure layout

- Choose two planes using the feature selection buttons

- Hit "Save Probe"

- PolyWorks will measure both planes

- Without creating any objects in the tree, you’ll see the distance between the planes

✅ Result

- Basic features like plane-to-plane distance captured without using full object tree

- Clean, fast, and beginner-friendly method to grab essential dimensions

📬 Further Support

- Once you're ready to go beyond the basics:

  - Contact: training@diversedimensions.com

  - Or call Diverse Dimensions directly

- The team can help elevate your PolyWorks proficiency

Using Thin Material Mode for Probing in PolyWorks

Presenter: Derek from Diverse Dimensions
Location: Eastside Lab – Brighton, Michigan
Focus: Probing thin sheet metal with FARO arm in PolyWorks

🎯 Purpose of the Feature

- When working with thin sheet metal, especially:

  - Scanned parts often handled with a scanner

  - But sometimes hard probing is still required

- PolyWorks offers a helpful option: Thin Material Mode

⚙️ How to Enable Thin Material Mode

- Go to the Properties of the feature you're creating

- Open the Measurement tab

- Check the box labeled “Thin Material”

- Enter the material thickness

🧠 What It Does

- Creates a “zone” for proper probing

- Ensures you:

  - Only collect points within valid thickness limits

  - Avoid points too far from ideal location

- Provides visual and audio feedback during probing:

  - 🔇 Invalid point = ignored or flagged

  - ✅ Valid click = point accepted

📊 Benefits

- Helps probe holes and features more accurately on thin parts

- Even if you’re slightly off the probe equator, PolyWorks:

  - Compensates using thin material calculations

- Prevents common errors when probing thin-walled geometry

🧪 Key Use Case

- Works well when hard probing must be used (even on scanned parts)

- Useful for quality control and reverse engineering on delicate sheet metal

📬 Need More Help?

- Reach out to Diverse Dimensions

- Contact for training or advanced techniques

- Don't forget to subscribe for more tutorials

Using Gauge Pins and Custom Measurements in PolyWorks

Presenter: Jeff Moss, Diverse Dimensions
Location: Eastside Lab, Brighton, MI
Focus: Adding gauge pin measurements alongside CMM/FARO arm data

🧰 Project Context

- Working with laser-cut tubing provided by a client

- Tubes contain holes that need to be checked by:

  - Digital measurement (FARO arm) for true position

  - Gauge pins for fit verification (per end-client request)

🎯 Goal

- Report both:

  - Digitally measured hole data

  - Physical gauge pin fit data

- Use custom measurements in PolyWorks to record manual inspection values

🧪 Measurement Details

- Example Hole: Dim 3-1

  - Digitally measured to ~25 microns of deviation

  - Laser cut and visually looks very clean

- Gauge pin:

  - Nominal size: 6 mm (0.2362 inches)

  - Fit test:

    - 0.232" gauge pin fit

    - 0.233" did not fit

  - Conclusion: 0.232" = verified pass size

🖥️ Using PolyWorks – Custom Measurements

- In PolyWorks 2020, go to:

  - Measure → Custom Measurements → Create

- Name the custom measurement:

  - Example: Dim 3-1 Gauge Pin

- Choose method: Numerical (manual entry)

- Input:

  - Nominal (CAD reference): 0.2362

  - Measured value: 0.232 (actual gauge pin used)

  - (Update tolerance later if needed)

- Result: Appears in model tree as a custom measurement

🧾 Why It’s Useful

- Keeps all data in one report:

  - CMM/FARO measured data

  - Gauge pin fit results

- Useful for:

  - True position + physical inspection feedback

  - Client reporting (clear and complete)

- Supports alphanumeric entries too:

  - e.g., “Yes”/“No” for pass/fail, threaded, visual inspection

Aligning Scan Data to CAD Using Best Fit & Point Pairs in PolyWorks

Presenter: Jeff Moss, Diverse Dimensions
Scenario: Scanning half of a mold and aligning point cloud to CAD
Objective: Demonstrate how to align scan (point cloud) data to CAD using point pairs and best fit

🧰 Setup Overview

- Project: Scanned one half of a mold for a client

- Captured:

  - All critical surfaces (mating interface)

  - Extra scan data on the side for orientation purposes

- Working in PolyWorks 2020, using:

  - Separate projects within the same workspace

  - Goal: Align point cloud to CAD model

📐 Alignment Method

Tool Used:

- Align → Best Fit → Data to Reference Objects

Steps:

- Open CAD model and scan data

- Use “Pre-alignment” option

  - Checked by default

  - Alignment method: Point Pairs

- Click Start

🖱️ Creating Point Pairs

- Two split screens appear for:

  - CAD model

  - Point cloud

- Rotate views to similar orientation

- Select matching points on both models in the same order:

  - 3 points on top surface

  - 2 points on flange

  - 1 point on end face
→ Total of 6 pairs

How to pick:

- Left-click to select points

- Cursor changes to pencil icon

- Ensure points are on surfaces that are mutually perpendicular (orthogonal), for better alignment

⚙️ Executing Best Fit

- After selecting point pairs:

  - Right-click to initiate alignment

  - PolyWorks performs an iterative alignment

  - Example:

    - Max iterations set to 50

    - PolyWorks converged in 6 iterations

Outcome:

- Point cloud is now aligned to CAD geometry

- Extra scan data still visible but can be cleaned before deviation or color map

🎯 Key Takeaways

- Use point pair alignment to get scan data roughly aligned

- Then use Best Fit to finalize precise registration

- Pick points that:

  - Are widely spaced

  - Represent multiple directions (X, Y, Z)

- Ensures accurate alignment for further surface deviation analysis and color mapping

🧠 Tips

- You don’t need full part coverage to align successfully

- Use clean features for alignment (e.g., flanges, planar faces, edges)

- Clean up excess scan data after alignment for better comparison results

Extracting Comparison Points from Point Cloud Data in PolyWorks

Presenter: Jeff Moss, Diverse Dimensions
Scenario: Scanned curved plywood cladding panels
Objective: Extract comparison points from a point cloud using defined criteria in PolyWorks

🧰 Project Context

- Scanned curved plywood panels

- Focused on the perimeter (border) of each panel

- Intention: Extract comparison points from scan data for analysis

- Used: PolyWorks v14, Scan to Matched alignment method

🖥️ Initial Setup

- Scan and best fit the point cloud to the CAD using:

  - A border feature around the panel

  - Followed by plane + axis + center point alignment

- Skinned mesh is now aligned to CAD surface

- Comparison points were placed using a sample grid:

  - Organizes points symmetrically

  - Keeps consistent spacing between points

📍 Configuring Comparison Points (Properties Setup)

Each point has specific settings in its properties:

- Radius:

  - Set to 0.250 inches (1/4")

  - Defines a zone around the nominal point where extraction is allowed

  - Like a tolerance bubble around the point

- Measurement Tab → Method:

  - Set to Extract from Point Cloud

- Parameters Tab:

  - Data object: Set to Surface

  - Max Distance: 0.5 inches

    - If deviation is > 0.5", point will not extract

    - Ensures outliers don’t get falsely mapped

🧪 How to Extract Points

- Select all the comparison points in the scene

- Right-click → Extract Measured

- PolyWorks will:

  - Find the nearest point from the point cloud within the defined radius and max distance

  - Color the points according to the deviation color map

    - Reflects how closely each scan point aligns to the nominal model

✅ Result

- Comparison points now have measured values tied to the scanned surface

- Deviation values visually reflected through color map

- Process is repeatable and scalable for large sets of points

🧠 Tips

- Use the sample grid tool (explained in another video) for structured placement

- Be sure to configure:

  - Extraction radius

  - Max deviation tolerance

- Check results with the color scale to interpret panel conformity